FILTERING  FACEPIECE RESPIRATORS

FFR   Care

	
	Table 2. Summary of the decontamination method and effect on FFR performance
	

	
	CDC
	
	
	
	
	

	
	Table 2 provides a summary of the decontamination methods evaluated in the referenced literature and the reported effect of each method on FFR performance.
	

	
	Method
	Treatment level
	FFR filtration performance
	FFR fit performance
	Other observations
	

	
	Vaporous hydrogen peroxide (VHP)
	Battelle report: Bioquell Clarus C HPV generator: The HPV cycle included a 10 min conditioning phase, 20 min gassing phase at 2 g/min, 150 min dwell phase at 0.5 g/min, and 300 min of aeration.
	Passed
	FFR fit was shown to be unaffected for up to 20 VHP treatments cycles using a head form
	Degradation of straps after 30 cycles (Battelle report)
	

	
	
	Bergman et. al.: Room Bio-Decontamination Service (RBDS™, BIOQUELL UK Ltd, Andover, UK), which utilizes four portable modules: the Clarus® R HPV generator (utilizing 30% H2O2), the Clarus R20 aeration unit, an instrumentation module and a control computer. Room concentration = 8 g/m3, 15 min dwell, 125 min total cycle time.
	
	
	
	

	
	Ultraviolet germicidal irradiation (UVGI)
	0.5–950 J/cm2
	Passed
	90–100% passing rate after 3 cycles depending on model
	
	

	
	Microwave generated steam
	1100–1250 W microwave models (range: 40 sec to 2 min)
	All models passed filtration evaluation for 1 or 20 treatment cycles as per test
	95–100% passing rate after 3 and 20 cycles for all models tested
	
	

	
	Microwave steam bags
	1100 W, 90 sec (bags filled with 60 mL tap water)
	Passed
	Not evaluated
	
	

	
	Moist heat incubation
	15 min–30 min (60°C, 80% RH)
	6 of 6 models passed after 3 cycles of contamination
	Passed
	
	

	
	Liquid hydrogen peroxide
	1 sec to 30 min (range: 3–6%)
	Passed
	Not evaluated
	
	

	
	Ethylene oxide
	1 hour at 55°C; conc. range: 725–833/L
	Passed
	Not evaluated
	
	


